## Sheet1
Vscan v6 -- APK Build : 2318 -- Device being checked on Android 9 -- Device Number VH263057X2
Functional & Adhoc Testing - 08-Nov-2021
DR ID | Test Steps performed | Observed result | Conclusion
1 | Functionality check , create new exam, edit patient. Check the image and video store on all presets. | Basic Functionality works as expected, image, videos gets stored and played / viewed. Both in BW mode and CF mode. | Works as expected.
2 | Standby check on scan screen | s301- (OS 9) - 2318 : stand by as expected - 10 mins +/- 1 mins | Works as expected.
s209 - (OS 5.1) - 2318 : auto freeze 10 mins, stand by as expected - 10 mins +/- 1 mins
3 | Standby check on dashboard screen | Standy by works as expected , in 10 minutes. | Works as expected.
4 | Standby check on settings screen | Standy by works as expected , in 10 minutes. | Works as expected.
5 | battery removal on scan screen | Removed battery on scan screen, the device gets shut down in 2 minutes. | Works as expected.
6 | Image servers/ worklist / tricefy / windows share | exams are shared successfully | Works as expected.
Battery Reinsert Scenario
| Screens | Result | 2318 final APK
| Protocol creator | Pass | Scan coach protocol crash popup displayed when battery reinserted on disclaimer screen and navigated to scan coach scan screen.
| Protocols | Pass
| protocol settings | Pass
| scan coach animation screen | Pass
| scan coach scan screen | Pass
| disclaimer screen. Steps overview | Pass
| edit patient | Pass
| menu | Pass
| worklist | Pass
| exam gallery | Pass
| settings | Pass
| about | Pass
| presets | Pass
| device status | Pass
| screen mirror | Pass
| server config | Pass
| admin settings | Pass
| android settings | Pass
| dashboard | Pass
| kiosk settings | Pass
| connect network | Pass
| GE Marketplace | Pass | Observation: While opening GE marketplace, the device asked to update the user credentials, and venkatesh's username and password was already available, however, it asked to confirm identity with the password, and without knowing the password, I exited and reopened GE Marketplace, it took me in without any identity confirmation and available apps were seen.
| Config | Pass
| View modules | Pass
| Enter a key (configuration) | Pass
| Remote update | Not Tested
| diagnostics | Pass | Observation: While checking for this test in system diagnostics, the device hangs and pauses the diagnostics and after Re-inserting the battery, the diagnostics wont resume and remains hanged. Have to manually restrat the device.
| display test | Pass
| Tricefy | Pass
| EA Uplink | Pass
| my account | Pass
| Lvivoef | Pass
| Bladder volume | Pass
| LUNG M. | Pass | Observation: while checking for this test in Lung M mode, after the battery reinsert, the scan screen exits from the LUNG M mode and redirects to a normal scan screen.